09/02/2026

Urgent Field Safety Notice

Additional warning concerning PENTAX Medical 
ultrasound endoscopes and water bottle assembly

Dear Healthcare Professional,

We are writing to inform you about a newly added warning for the Instructions for Use (IFU). This update is part of PENTAX’ ongoing commitment to patient safety.

- Background

PENTAX Medical has received an information related to a reported patient death in connection with a case in which air unintentionally entered the balloon attached to the distal end of an ultrasound endoscope during an endoscopic ultrasonography procedure.

As a result of the investigation, a definitive causal relationship between the ingress of air into the balloon and the patient’s death could not be established. However, it also cannot be completely ruled out. Therefore, to clearly address this potential risk, PENTAX Medical has added a new warning to the IFU for the relevant ultrasound endoscopes and the water bottle assembly. This additional warning will be included in the IFU of all newly manufactured devices. For devices already placed on the market the users shall add this Field Safety Notice to their documentation and inform their staff accordingly.

As this issue may, under certain circumstances, potentially lead to a serious adverse event, PENTAX Medical provides the following information to all customers using ultrasound endoscopes to ensure continued safe use of our products.

- Additional warning update

The following warning is added to the IFU of PENTAX Medical products.

<Ultrasound endoscopes>

ALWAYS ensure there is deaerated sterile water in the water bottle assembly prior to and during the procedure. Performing ultrasound balloon water feeding with an empty water bottle assembly can result in the balloon filling with air. When the ultrasound balloon fills with air, the ultrasound image will become unclear and/or disappear entirely. Furthermore, if feeding is continued with an empty water bottle assembly without notice of this condition, it may cause the balloon to rupture and may pose a serious risk to the patient. Risks to the patient include unintended pain and, in some cases, perforation of the gastrointestinal tract which may require medical intervention and in severe cases it may result in death.

<Water bottle assembly>

ALWAYS ensure there is deaerated sterile water in the water bottle assembly prior to and during the procedure when performing the balloon method together with the ultrasound upper GI videoscope. Performing ultrasound balloon water feeding with an empty water bottle assembly can result in the balloon filling with air. When the ultrasound balloon fills with air, the ultrasound image will become unclear and/or disappear entirely. Furthermore, if feeding is continued with an empty water bottle assembly without notice of this condition, it may cause the balloon to rupture and may pose a serious risk to the patient. Risks to the patient include unintended pain and, in some cases, perforation of the gastrointestinal tract which may require medical intervention and in severe cases it may result in death.

- Affected medical devices

The additional warning is updated to the IFUs of the products listed in the table below.

| Category | Model name | Serial number(s) Lot numbers(s) | UDI-DI |
| --- | --- | --- | --- |
| Ultrasound endoscope | EG36-J10UR | All | 04961333239573 04961333239580 04961333239597 04961333239603 04961333239627 04961333248919 04961333248926 04961333248933 04961333248957 04961333252756 |
|  | EG38-J10UT | All | 04961333239474 04961333239481 04961333239498 04961333239504 04961333239528 04961333252770 |
|  | EG34-J10U | All | 04961333239672 04961333239689 04961333239696 04961333239702 04961333239726 04961333248834 04961333248841 04961333248858 04961333248872 04961333248896 04961333252732 |
|  | EG-3870UTK | All | 04961333237098 04961333239344 04961333226603 04961333228409 04961333254217 04961333237104 04961333225828 04961333228423 |
|  | EG-3670URK | All | 04961333237074 04961333237081 04961333226597 04961333228362 04961333254200 04961333228386 |
|  | EG-3270UK | All | 04961333237050 04961333237067 04961333226580 04961333254187 04961333233793 |
| Water bottle assembly | OS-H5 | All | 04961333219940 04961333247486 04961333251032 04961333251049 |

Actions to be taken by the user

To avoid the ingress of air into the balloon in place of water during ultrasound endoscopic procedures, the following steps must be strictly followed. These steps are already described in current IFUs of PENTAX Medical ultrasound endoscopes and water bottle assembly.

- Users should inform their staff accordingly of the additional warning and updates to IFU.

- Change water for each procedure (in IFU of Water bottle assembly).

- Fill the bottle approximately 2 / 3 full (to the “200” marking on the bottle scale) with sterile water (in IFU of Water bottle assembly).

- If “The water/ balloon feeding function does not work”, please “Fill the water bottle with deaerated sterile water” (in IFU of ultrasound endoscope).

- Contact Information:

PENTAX Medical regrets any inconvenience that this action may cause and appreciates your understanding and cooperation. Please be assured that maintaining product quality is our utmost priority.

Please indicate through the attached response form that you have received and understood this information, by completing it and returning no later than 30 days from receipt.

it by Sep. 30th, 2026, at: customeradvisories@pentaxmedical.com

We look forward to your continued patronage of PENTAX Medical products. If you have any questions regarding this notice, please feel free to contact us at:

- Tel: 1-800-431-5880 (8:30 AM – 5:00 PM, Monday – Friday, EST)

- Email: customeradvisories@pentaxmedical.com

Adverse events / serious incidents experienced with the use of this product should be reported as soon as possible to PENTAX Medical at vigilance@pentaxmedical.com.

Sincerely,

Motoki TOGASHI

Chief Quality Officer

PENTAX Medical

CUSTOMER RESPONSE FORM
RESPONSE IS REQUIRED

Ref: 2026-002-C / FSCA-PMJ-26-03-2 Rev.00

Field Safety Notice for Request for IFU updates for multiple ultrasound endoscopes and water bottle assembly

Please complete this form and return it to PENTAX Medical promptly upon receipt and no later than 30 days from receipt. This will confirm receipt and understanding of the Field Safety Notice.

Facility Name:        __________________________________________________

Street Address:       __________________________________________________

City/State/ZIP/Country: __________________________________________________

By signing this form, we acknowledge receipt and understanding of the accompanying Field Safety Notification, Downloaded the applicable IFU (https://ifu.pentaxmedical.com) and that we have informed all potential users and have taken and will take appropriate actions in accordance with that Notification.

If you requested the additional training for this notification, please check the “Yes” box below.

☐ Yes

☐ We no longer own these equipment/s.

Please provide the name of the individual with responsibility who completed this form.

Signature:               	______________________________________________

Name:                  	______________________________________________

Customer Phone Number:	______________________________________________

Customer Email Address: 	______________________________________________

Date:                  	______________________________________________

Upon completion of the form and signing, please return the form:

- Email a pdf copy of the completed form to customeradvisories@pentaxmedical.com

If you have any questions regarding this action, please feel free to contact your PENTAX Medical Territory Manager or PENTAX Medical Customer Service at 1-800-431-5880.